## Sheet 1
Laboratory | Facility | Country | City | Sample Population | Sample Sub-Population | Reference Sample Type | Sample Set | Reference Method | Repeats | Blinding | Test Type | Test Format | Use Case | Test Name | Manufacturer | Catalog Number | Lot Number | Index Sample Type | Target | Antigen Used for Detection | RNA Extraction Kit | Estimated LoD | LoD Unit | Higher Cut-off | Lower Cut-off | Ct Value | Days of Sickness | Sample Size | Total Positive | Total Negative | True Positive | True Negative | Sensitivity | Specificity | Link to Publication | DOI | Comments
Servicio de Laboratorio Clínico, Hospital Clínico de la Universidad de Chile "Dr. José Joaquín Aguirre" | Hospital | Chile | Santiago | Confirmed COVID-19 patients | Not available | Nasopharyngeal swabs | Retrospective set | RT-PCR | No replicates | Not blinded to reference results | Molecular | Manual PCR | Not available | 2019-nCoV TaqMan RT-PCR Kit (Catalog# TM67100) | Thermo Fisher Scientific | Not available | Not available | Nasopharyngeal swabs | Specimen without RNA extraction | Not available | Not available | Not available | Not available | Not available | Not available | Not available | Overall | 2 | 2 | 0 | 2 | 0 | 100 |  | https://www.biorxiv.org/content/10.1101/2020.03.28.013508v1.full.pdf | 10.1101/2020.03.28.013508 | Not available
Mammoth Biosciences, Inc., San Francisco, California, USA And Department of Laboratory Medicine, University of California, San Francisco, California, USA | Public Health | United States of America (USA) | Not available | Confirmed COVID-19 patients | Not available | Nasopharyngeal swabs; Oropharyngeal swabs | Prospectively collected set | RT-PCR | No replicates | Not blinded to reference results | Molecular | CRISPR-based lateral flow assay | Not available | SARS-CoV-2 DETECTR | Mammoth Biosciences | Not available | Not available | Nasopharyngeal swabs; Oropharyngeal swabs | Fresh RNA extracted | Not available | Qiagen DSP Viral RNA Mini Kit | 10 | copy / microL | Not available | Not available | Not available | Overall | 21 | 10 | 11 | 9 | 11 | 90 | 100 | https://www.medrxiv.org/content/10.1101/2020.03.06.20032334v2.full.pdf | 10.1101/2020.03.06.20032334 | Note, manufacturer is among authors.
National Institute for Communicable Diseases (NICD) | Public Health | South Africa | Not available | Not available | Not available | Nasopharyngeal swabs | Not available | RT-PCR | No replicates | Not available | Molecular | Manual PCR | Not available | COVID-19 Nucleic Acid Detection Kit (Multiplex Real Time PCR) | Sansure Biotech, Inc. | Not available | Not available | Nasopharyngeal swabs | Not available | Not available | MagnaPure 96 | 0.1 | copy / microL | Not available | Not available | Not available | Overall | 21 | 16 | 5 | 15 | 4 | 93.8 | 80 | Not available | Not available | Not available
National Institute for Communicable Diseases (NICD) | Public Health | South Africa | Not available | Not available | Not available | Nasopharyngeal swabs | Not available | RT-PCR | No replicates | Not available | Molecular | Manual PCR | Not available | LightMix Modular SARS-CoV-2 (COVID19) | TIB Molbiol/Roche Diagnostics | Not available | Not available | Nasopharyngeal swabs | Not available | Not available | MagnaPure 96 | 0 | copy / microL | Not available | Not available | Not available | Overall | 27 | 20 | 7 | 16 | 7 | 80 | 100 | Not available | Not available | LoD: 0.01 copies/uM Egene; 0.1 copies/uM Ngene; 0.15 copies/uM RdRP
National Institute for Communicable Diseases (NICD) | Public Health | South Africa | Not available | Not available | Not available | Nasopharyngeal swabs | Not available | RT-PCR | No replicates | Not available | Molecular | Manual PCR | Not available | Real-time fluorescent RT-PCR kit for detecting 2019 nCoV | BGI Health (HK) Co. Ltd. | Not available | Not available | Nasopharyngeal swabs | Not available | Not available | MagnaPure 96 | 0.1 | copy / microL | Not available | Not available | Not available | Overall | 27 | 20 | 7 | 14 | 7 | 70 | 100 | Not available | Not available | LoD: 0.1 copies/uL on Ngene; 0.15 copies/uL on RdRP; 1.3 c/uL Egene
National Institute for Communicable Diseases (NICD) | Public Health | South Africa | Not available | Not available | Not available | Nasopharyngeal swabs | Not available | RT-PCR | No replicates | Not available | Molecular | Manual PCR | Not available | YouSeq SARS-COV-2 qPCR Test | YouSeq Ltd | Not available | Not available | Nasopharyngeal swabs | Not available | Not available | MagnaPure 96 | 0 | copy / microL | Not available | Not available | Not available | Overall | 21 | 16 | 5 | 13 | 5 | 81.2 | 100 | Not available | Not available | LoD determined on Egene
National Institute for Communicable Diseases (NICD) | Public Health | South Africa | Not available | Not available | Not available | Nasopharyngeal swabs | Not available | RT-PCR | No replicates | Not available | Molecular | Manual PCR | Not available | TaqPath COVID-19 Combo Kit | Thermo Fisher Scientific | Not available | Not available | Nasopharyngeal swabs | Not available | Not available | MagnaPure 96 | 1.4 | copy / microL | Not available | Not available | Not available | Overall | 33 | 26 | 7 | 13 | 7 | 50 | 100 | Not available | Not available | LoD determined on Ngene
National Institute for Communicable Diseases (NICD) | Public Health | South Africa | Not available | Not available | Not available | Nasopharyngeal swabs | Not available | RT-PCR | No replicates | Not available | Molecular | Manual PCR | Not available | Logix Smart Coronavirus disease 2019 (COVID-19) | Co-diagnostics | Not available | Not available | Nasopharyngeal swabs | Not available | Not available | MagnaPure 96 | 0.2 | copy / microL | Not available | Not available | Not available | Overall | Not available | 6 | Not available | 4 | Not available | 66.7 |  | Not available | Not available | LoD determined on RdRP
National Institute for Communicable Diseases (NICD) | Public Health | South Africa | Not available | Not available | Not available | Nasopharyngeal swabs | Not available | RT-PCR | No replicates | Not available | Molecular | Manual PCR | Not available | Bosphore Novel Coronavirus (2019-nCoV) Detection Kit | Anatolia Geneworks | Not available | Not available | Nasopharyngeal swabs | Not available | Not available | MagnaPure 96 | Not available | Not available | Not available | Not available | Not available | Overall | Not available | 6 | Not available | 3 | Not available | 50 |  | Not available | Not available | LoD: below 13 copies/uL
UCSD | Hospital | United States of America (USA) | San Diego | Symptomatic patients | Not available | Nasopharyngeal swabs | Prospectively collected set | RT-PCR | Not available | Blinded to reference results | Molecular | Integrated system | Not available | Fluxergy Sample-to-Answer SARS-CoV-2 RT-PCR | Fluxergy LLC | Not available | Not available | Nasopharyngeal swabs | Specimen without RNA extraction | Not available | Not available | 33 | copy / reaction | Not available | Not available | Not available | Overall | 92 | 31 | 61 | 29 | 58 | 93.5 | 95.1 | Not available | Not available | Not available
Second People's Hospital of Fuyang | Hospital | China | Fuyang | Confirmed COVID-19 patients | Mild; Severe; Critical | Not available | Retrospective set | Clinical examination; Computed Tomography (CT); RT-PCR | Not available | Not available | Molecular | Not available | Not available | Not available | Shanghai BioGerm Medical Biotechnology Co., Ltd. | Not available | Not available | Oropharyngeal swabs | Not available | Not available | Viral nucleic acid isolation kit from Jiangsu Bioperfectus Technologies Company, Ltd. | Not available | Not available | Not available | Not available | Not available | Overall | 38 | 38 | 0 | 14 | Not available | 36.8 |  | https://www.medrxiv.org/content/10.1101/2020.03.26.20042044v1 | 10.1101/2020.03.26.20042044 | Not available
Second People's Hospital of Fuyang | Hospital | China | Fuyang | Confirmed COVID-19 patients | Mild; Severe; Critical | Not available | Retrospective set | Clinical examination; Computed Tomography (CT); RT-PCR | Not available | Not available | Molecular | Not available | Not available | Not available | Shanghai BioGerm Medical Biotechnology Co., Ltd. | Not available | Not available | Sputum | Not available | Not available | Viral nucleic acid isolation kit from Jiangsu Bioperfectus Technologies Company, Ltd. | Not available | Not available | Not available | Not available | Not available | Overall | 38 | 38 | 0 | 29 | Not available | 76.3 |  | https://www.medrxiv.org/content/10.1101/2020.03.26.20042044v1 | 10.1101/2020.03.26.20042044 | Not available
Microbiology Department, Basingstoke & North Hampshire Hospital | Hospital | Great Britain | Not available | Not available | Not available | Nasopharyngeal swabs; Oropharyngeal swabs | Not available | RT-PCR | Not available | Not available | Molecular | Manual PCR | Not available | genesig Real-Time PCR COVID-19 | Novacyt/primerdesign | Not available | Not available | Nasopharyngeal swabs; Oropharyngeal swabs | Not available | Not available | PerkinElmer chemagic Prepito-D (NA Body Fluid kit; plasma protocol) | Not available | Not available | Not available | Not available | 28.3 | Overall | 57 | 4 | 53 | 4 | 53 | 100 | 100 | Not available | Not available | Target: RdRP gene target
Microbiology Department, Basingstoke & North Hampshire Hospital | Hospital | Great Britain | Not available | Not available | Not available | Nasopharyngeal swabs; Oropharyngeal swabs | Not available | RT-PCR | Not available | Not available | Molecular | Manual PCR | Not available | VIASURE SARS-CoV-2 S gene Real Time PCR Detection Kit | CerTest Biotec, S.L. | Not available | Not available | Nasopharyngeal swabs; Oropharyngeal swabs | Not available | Not available | PerkinElmer chemagic Prepito-D (NA Body Fluid kit; plasma protocol) | Not available | Not available | Not available | Not available | 28.6 | Overall | 29 | 0 | 29 | Not available | 29 |  | 100 | Not available | Not available | VIASURE 2019-nCoV – S gene target.
NIS Laboratories, National Infection Service, Public Health England | Public Health | Great Britain | Not available | Not available | Not available | Respiratory specimen | Not available | RT-PCR | Not available | Not available | Molecular | Manual PCR | Not available | RealStar® SARS-CoV-2 RT-PCR Kit | altona Diagnostics | Not available | Not available | Respiratory specimen | Not available | Not available | Not available | Not available | Not available | Not available | Not available | Not available | Overall | 195 | 0 | 195 | Not available | 195 |  | 100 | Not available | Not available | Target: B-ß coronavirus RNA
NIS Laboratories, National Infection Service, Public Health England | Public Health | Great Britain | Not available | Not available | Not available | Respiratory specimen | Not available | RT-PCR | Not available | Not available | Molecular | Integrated system | Not available | SARS-CoV-2, Influenza and RSV 8-well | AusDiagnostics | #20619 | Not available | Respiratory specimen | Not available | Not available | Not available | Not available | Not available | Not available | Not available | Not available | Overall | 195 | 0 | 195 | Not available | 195 |  | 100 | Not available | Not available | Target: SARS-CoV-2 Orf1ab
NIS Laboratories, National Infection Service, Public Health England | Public Health | Great Britain | Not available | Not available | Not available | Nasopharyngeal swabs; Oropharyngeal swabs; Respiratory specimen | Not available | RT-PCR | Not available | Not available | Molecular | Integrated system | Not available | EasyScreenTM Pan-Coronavirus/SARS-CoV-2 Detection Kit | Genetic Signatures Limited | RP010-HT | Not available | Respiratory specimen | Not available | Not available | Not available | Not available | Not available | Not available | Not available | Not available | Overall | 195 | 0 | 195 | Not available | 195 |  | 100 | Not available | Not available | Target: E-gene
National Infection Service, Public Health England | Public Health | Great Britain | Not available | Not available | Not available | Respiratory specimen | Not available | RT-PCR | Not available | Not available | Molecular | Integrated system | Not available | Extended Coronavirus Array (lab-based or near-POC) | Randox Laboratories Ltd | EV4418 | Not available | Respiratory specimen | Not available | Not available | Not available | Not available | Not available | Not available | Not available | Not available | Overall | 195 | 0 | 195 | Not available | 195 |  | 100 | Not available | Not available | Not available
Department of Virology, Manchester Medical Microbiology Partnership, Manchester Royal Infirmary | Public Health | Great Britain | Manchester | Not available | Not available | Nasopharyngeal swabs; Oropharyngeal swabs; Sputum | Not available | RT-PCR | Not available | Not available | Molecular | Integrated system | Not available | Coronavirus LightMix® Modular SARS and Wuhan CoV E-gene assay | Roche Ltd | 09155368001 | Not available | Nasopharyngeal swabs; Oropharyngeal swabs; Sputum | Not available | Not available | Magna Pure 96 using the Pathogen Universal protocol 200 v4.0 | Not available | Not available | Not available | Not available | Not available | Overall | 165 | 15 | 150 | 15 | 150 | 100 | 100 | Not available | Not available | Not available
National Infection Service, Public Health England | Public Health | Great Britain | Not available | Not available | Not available | Respiratory specimen | Not available | RT-PCR | Not available | Not available | Molecular | Manual PCR | Not available | The Novel Coronavirus (2019-nCoV) Nucleic Acid Test Kit | GeneFirst Ltd | Not available | Not available | Respiratory specimen | Not available | Not available | Not available | Not available | Not available | Not available | Not available | Not available | Overall | 195 | 0 | 195 | Not available | 195 |  | 100 | Not available | Not available | Target: orf1ab
NIS Laboratories, National Infection Service, Public Health England | Public Health | Great Britain | Not available | Not available | Not available | Respiratory specimen | Not available | RT-PCR | Not available | Not available | Molecular | Manual PCR | Not available | genesig Real-Time PCR COVID-19 | Novacyt/primerdesign | Z-Path-COVID-19-CE | Not available | Respiratory specimen | Not available | Not available | Not available | Not available | Not available | Not available | Not available | Not available | Overall | 195 | 0 | 195 | Not available | 195 |  | 100 | Not available | Not available | Target: RdRp of SARS-CoV-2
NIS Laboratories, National Infection Service, Public Health England | Public Health | Great Britain | Not available | Not available | Not available | Respiratory specimen | Not available | RT-PCR | Duplicates | Not available | Molecular | Manual PCR | Not available | Allplex 2019-nCoV assay | Seegene, Inc. | #RP10244Y | Not available | Respiratory specimen | Not available | Not available | Not available | Not available | Not available | Not available | Not available | Not available | Overall | 195 | 0 | 195 | Not available | 195 |  | 100 | Not available | Not available | Target: SARS-CoV-2 RdRP and N genes
Wuhan Jinyintan Hospital | Hospital | China | Wuhan | Confirmed COVID-19 patients | Not available | Not available | Not available | WHO Interim Guidance | Not available | Not available | Molecular | Manual PCR | Not available | Liferiver Novel Coronavirus (2019-nCoV) Real Time Multiplex RT-PCRT kit | Shanghai ZJ Bio-Tech Co., Ltd/Liferiver | Not available | Not available | Nasopharyngeal swabs; Oropharyngeal swabs; Sputum | Fresh RNA extracted | Not available | Qiamp® viral RNA mini kit | Not available | Not available | Not available | Not available | Not available | Overall | Not available | 81 | Not available | 42 | Not available | 51.9 |  | https://www.medrxiv.org/content/10.1101/2020.02.27.20028787v1.full.pdf+html | 10.1101/2020.02.27.20028787 | Not available
Renmin Hospital of Wuhan University | Hospital | China | Wuhan | Confirmed COVID-19 patients | Moderate; Severe; Critical | Not available | Retrospective set | Clinical examination; Computed Tomography (CT); RT-PCR | Not available | Not available | Molecular | Manual PCR | Not available | SARS-CoV-2 ORF1ab/N qPCR detection kit | Shanghai GeneoDx Biotechnology Co., Ltd | Not available | Not available | Not available | Not available | Not available | Not available | Not available | Not available | Not available | Not available | Not available | Overall | 133 | 133 | 0 | 91 | Not available | 68.4 |  | https://www.medrxiv.org/content/10.1101/2020.03.28.20045765v1 | 10.1101/2020.03.28.20045765 | Not available
Centre National de Référence (CNR) des virus des infections respiratoires | Public Health | France | Lyon | Confirmed COVID-19 patients | Not available | Nasopharyngeal swabs; Deep respiratory specimen | Retrospective set | RT-PCR | No replicates | Not available | Molecular | Integrated system | Not available | Xpert Xpress SARS-CoV-2 | Cepheid | 900-0883 | Not available | Nasopharyngeal swabs; Deep respiratory specimen | Specimen without RNA extraction | Not available | Not available | Not available | Not available | Not available | Not available | 39 | Overall | 8 | 5 | 3 | 5 | 3 | 100 | 100 | https://www.sfm-microbiologie.org/wp-content/uploads/2020/04/RESULTATS-CNR-CEPHEID-tests-RUO-_-CNR-LYON-_30032020.pdf | Not available | Not available
Centre National de Référence (CNR) des virus des infections respiratoires | Public Health | France | Lyon | Confirmed COVID-19 patients | Not available | Nasopharyngeal swabs; Deep respiratory specimen | Retrospective set | RT-PCR | No replicates | Not available | Molecular | Integrated system | Not available | BioGX SARS-CoV-2 Open System Reagents for BD MAX™ | BioGX | 350-093-C-MAX | Not available | Nasopharyngeal swabs; Deep respiratory specimen | Specimen without RNA extraction | Not available | Not available | Not available | Not available | Not available | Not available | 40.3 | Overall | 10 | 7 | 3 | 7 | 3 | 100 | 100 | https://www.sfm-microbiologie.org/2020/04/03/covid-19/ | Not available | Not available
Centre National de Référence (CNR) des virus des infections respiratoires | Public Health | France | Lyon | Confirmed COVID-19 patients | Not available | Nasopharyngeal swabs; Deep respiratory specimen | Retrospective set | RT-PCR | No replicates | Not available | Molecular | Integrated system | Not available | SARS-COV-2 R-GENE® (ref 423717) | bioMérieux SA | 423717 | Not available | Nasopharyngeal swabs; Deep respiratory specimen | Specimen without RNA extraction | Not available | Not available | Not available | Not available | Not available | Not available | 35 | Overall | 9 | 6 | 3 | 6 | 3 | 100 | 100 | https://www.sfm-microbiologie.org/2020/04/03/covid-19/ | Not available | Not available
Centre National de Référence (CNR) des virus des infections respiratoires | Public Health | France | Lyon | Confirmed COVID-19 patients | Not available | Respiratory specimen | Retrospective set | RT-PCR | No replicates | Not available | Molecular | Integrated system | Not available | cobas® SARS-CoV-2 (for use on the cobas® 6800/8800 Systems) | Roche Molecular Diagnostics | 09175431190 | Not available | Respiratory specimen | Specimen without RNA extraction | Not available | Not available | Not available | Not available | Not available | Not available | 35 | Overall | 20 | 15 | 5 | 15 | 5 | 100 | 100 | https://www.sfm-microbiologie.org/2020/04/03/covid-19/ | Not available | Not available
Centre National de Référence (CNR) des virus des infections respiratoires | Public Health | France | Lyon | Not available | Not available | Nasopharyngeal swabs; Deep respiratory specimen | Prospectively collected set | RT-PCR | No replicates | Not available | Molecular | Integrated system | Not available | VitaPCR™ SARS-CoV-2 Assay | BIOSYNEX S.A. | Not available | Not available | Respiratory specimen | Specimen without RNA extraction | Not available | Not available | Not available | Not available | Not available | Not available | 26 | Overall | 15 | 8 | 7 | 7 | 7 | 87.5 | 100 | https://www.sfm-microbiologie.org/2020/04/03/covid-19/ | Not available | Not available
Centre National de Référence (CNR) des virus des infections respiratoires | Public Health | France | Lyon | Confirmed COVID-19 patients | Not available | Nasopharyngeal swabs; Deep respiratory specimen | Retrospective set | RT-PCR | No replicates | Not available | Molecular | Manual PCR | Not available | VIASURE SARS-CoV-2 S gene Real Time PCR Detection Kit adapted for BD MAX™ System | CerTest Biotec, S.L. | Not available | Not available | Nasopharyngeal swabs; Deep respiratory specimen | Specimen without RNA extraction | Not available | Not available | Not available | Not available | Not available | Not available | 35 | Overall | 10 | 7 | 3 | 5 | 3 | 71.4 | 100 | https://www.sfm-microbiologie.org/2020/04/03/covid-19/ | Not available | Not available
Centre National de Référence (CNR) des virus des infections respiratoires | Public Health | France | Lyon | Confirmed COVID-19 patients | Not available | Nasopharyngeal swabs; Deep respiratory specimen | Retrospective set | RT-PCR | No replicates | Not available | Molecular | Integrated system | Not available | VitaPCR™ SARS-CoV-2 Assay | BIOSYNEX S.A. | Not available | Not available | Nasopharyngeal swabs; Deep respiratory specimen | Specimen without RNA extraction | Not available | Not available | Not available | Not available | Not available | Not available | Not available | Overall | 11 | 8 | 3 | 3 | 3 | 37.5 | 100 | https://www.sfm-microbiologie.org/2020/04/03/covid-19/ | Not available | Not available
Centre National de Référence (CNR) des virus des infections respiratoires | Public Health | France | Lyon | Confirmed COVID-19 patients | Not available | Nasopharyngeal swabs; Deep respiratory specimen | Retrospective set | RT-PCR | No replicates | Not available | Molecular | Manual PCR | Not available | RealStar® SARS-CoV-2 RT-PCR Kit | altona Diagnostics | 821003 | Not available | Nasopharyngeal swabs; Deep respiratory specimen | Fresh RNA extracted | Not available | Not available | Not available | Not available | Not available | Not available | 35 | Overall | 10 | 8 | 2 | 7 | 2 | 87.5 | 100 | https://www.sfm-microbiologie.org/2020/04/03/covid-19/ | Not available | Not available
Hopital Charles Le Moyne | Hospital | Canada | Not available | Symptomatic patients | Mild; Moderate; Severe; Critical | Nasopharyngeal swabs; Oropharyngeal swabs | Prospectively collected set | RT-PCR | No replicates | Blinded to reference results | Molecular | Manual PCR | Not available | RIDA® GENE SARS-CoV-2 RUO (PG6815RUO) | R-Biopharm AG | Not available | Not available | Nasopharyngeal swabs; Oropharyngeal swabs | Specimen without RNA extraction | Not available | Not available | Not available | Not available | Not available | Not available | Not available | Overall | 20 | 5 | 15 | 5 | 15 | 100 | 100 | Not available | Not available | Not available
Institut Pasteur | Private Laboratory | France | Paris | Confirmed COVID-19 patients | Not available | Nasopharyngeal swabs | Retrospective set | RT-PCR | High concentration pools: tested once; Low/neg concentration pools: tested in triplicate | Not available | Molecular | Manual PCR | Not available | TaqMan™ 2019-nCoV Assay Kit v1 | Thermo Fisher Scientific | Not available | Not available | Nasopharyngeal swabs | Frozen RNA | Not available | Extraction NucleoSpin Dx Virus (Réf. Macherey Nagel 740895.50) | Not available | Not available | Not available | Not available | Not available | Overall | 9 | 8 | 1 | 8 | 1 | 100 | 100 | https://www.sfm-microbiologie.org/wp-content/uploads/2020/04/RESULTATS-Thermo-v1.pdf | Not available | Samples tested were a serial dilution pool of samples.
Institut Pasteur | Private Laboratory | France | Paris | Confirmed COVID-19 patients | Not available | Nasopharyngeal swabs | Retrospective set | RT-PCR | High concentration pools: tested once; Low/neg concentration pools: tested in triplicate | Not available | Molecular | Manual PCR | Not available | TaqPath COVID-19 Combo Kit | Thermo Fisher Scientific | Not available | Not available | Nasopharyngeal swabs | Frozen RNA | Not available | Extraction NucleoSpin Dx Virus (Réf. Macherey Nagel 740895.50) | Not available | Not available | Not available | Not available | Not available | Overall | 9 | 8 | 1 | 8 | 1 | 100 | 100 | https://www.sfm-microbiologie.org/wp-content/uploads/2020/04/RESULTATS-Thermo-v2.pdf | Not available | Samples tested were a serial dilution pool of samples; performance when amplified on QuantStudio 5/6/7/12
Institut Pasteur | Private Laboratory | France | Paris | Confirmed COVID-19 patients | Not available | Nasopharyngeal swabs | Retrospective set | RT-PCR | High concentration pools: tested once; Low/neg concentration pools: tested in triplicate | Not available | Molecular | Manual PCR | Not available | RIDA® GENE SARS-CoV-2 RUO (PG6815RUO) | R-Biopharm AG | Not available | Not available | Nasopharyngeal swabs | Frozen RNA | Not available | Extraction NucleoSpin Dx Virus (Réf. Macherey Nagel 740895.50) | Not available | Not available | Not available | Not available | Not available | Overall | 9 | 8 | 1 | 8 | 1 | 100 | 100 | https://www.sfm-microbiologie.org/wp-content/uploads/2020/04/RESULTATS-RIDAgene.pdf | Not available | Samples tested were a serial dilution pool of samples; performance when amplified on ABI 7500 Fast
Institut Pasteur | Private Laboratory | France | Paris | Confirmed COVID-19 patients | Not available | Nasopharyngeal swabs | Retrospective set | RT-PCR | High concentration pools: tested once; Low/neg concentration pools: tested in triplicate | Not available | Molecular | Manual PCR | Not available | RIDA® GENE SARS-CoV-2 RUO (PG6815RUO) | R-Biopharm AG | Not available | Not available | Nasopharyngeal swabs | Frozen RNA | Not available | Extraction NucleoSpin Dx Virus (Réf. Macherey Nagel 740895.50) | Not available | Not available | Not available | Not available | Not available | Overall | 9 | 8 | 1 | 7 | 1 | 87.5 | 100 | https://www.sfm-microbiologie.org/wp-content/uploads/2020/04/RESULTATS-RIDAgene.pdf | Not available | Samples tested were a serial dilution pool of samples; performance when amplified on LightCycler 480
Institut Pasteur | Private Laboratory | France | Paris | Confirmed COVID-19 patients | Not available | Nasopharyngeal swabs | Retrospective set | RT-PCR | High concentration pools: tested once; Low/neg concentration pools: tested in triplicate | Not available | Molecular | Manual PCR | Not available | genesig Real-Time PCR COVID-19 | Novacyt/primerdesign | Not available | Not available | Nasopharyngeal swabs | Frozen RNA | Not available | Extraction NucleoSpin Dx Virus (Réf. Macherey Nagel 740895.50) | Not available | Not available | Not available | Not available | Not available | Overall | 9 | 8 | 1 | 8 | 1 | 100 | 100 | https://www.sfm-microbiologie.org/wp-content/uploads/2020/04/RESULTATS-genesig.pdf | Not available | Serial dilution of pooled samples
Institut Pasteur | Private Laboratory | France | Paris | Confirmed COVID-19 patients | Not available | Nasopharyngeal swabs | Retrospective set | RT-PCR | High concentration pools: tested once; Low/neg concentration pools: tested in triplicate | Not available | Molecular | Manual PCR | Not available | TaqPath COVID-19 Combo Kit | Thermo Fisher Scientific | Not available | Not available | Nasopharyngeal swabs | Frozen RNA | Not available | Extraction NucleoSpin Dx Virus (Réf. Macherey Nagel 740895.50) | Not available | Not available | Not available | Not available | Not available | Overall | 9 | 8 | 1 | 7 | 1 | 87.5 | 100 | https://www.sfm-microbiologie.org/wp-content/uploads/2020/04/RESULTATS-Thermo-v2.pdf | Not available | Samples tested were a serial dilution pool of samples; performance when amplified on LightCycler 480/ABI 7500 Fast
Nantong Third Hospital Affiliated to Nantong University, Clinical Laboratory | Public Health | China | Not available | Symptomatic patients | Not available | Pharyngeal swabs | Prospectively collected set | RT-PCR | Not available | Not available | Molecular | Integrated system | Not available | RainSure COVID-19 dPCR Detection Kit (lab-based) | RainSure Scientific Co., Ltd | Not available | Not available | Pharyngeal swabs | Not available | Not available | Not available | Not available | Not available | Not available | Not available | Not available | Overall | 108 | 28 | 80 | 25 | 76 | 89.3 | 95 | https://www.medrxiv.org/content/10.1101/2020.03.24.20042689v1.full.pdf | 10.1101/2020.03.24.20042689 | Target gene: ORF1ab
Centre National de Référence (CNR) Virus des Infections Respiratoires | Public Health | France | Paris | Confirmed COVID-19 patients | Not available | Nasopharyngeal swabs | Retrospective set | RT-PCR | High concentration pools: tested once; Low/neg concentration pools: tested in triplicate | Not available | Molecular | Integrated system | Not available | Amplidiag COVID-19 Kit | Mobidiag | Not available | Not available | Nasopharyngeal swabs | Specimen without RNA extraction | Not available | Not available | Not available | Not available | Not available | Not available | 39.4 | Overall | 10 | 8 | 2 | 8 | 2 | 100 | 100 | https://www.sfm-microbiologie.org/2020/04/03/covid-19/ | Not available | Ct value given is for weakest dilution pool. Amplification with BioRad CFX96
Centre National de Référence (CNR) Virus des Infections Respiratoires | Public Health | France | Paris | Confirmed COVID-19 patients | Not available | Respiratory specimen | Retrospective set | RT-PCR | Not available | Not available | Molecular | Manual PCR | Not available | The Novel Coronavirus (2019-nCoV) Nucleic Acid Test Kit | GeneFirst Ltd | Not available | Not available | Respiratory specimen | Frozen RNA | Not available | Not available | Not available | Not available | Not available | Not available | 36.7 | Overall | 10 | 8 | 2 | 8 | 2 | 100 | 100 | https://www.sfm-microbiologie.org/2020/04/03/covid-19/ | Not available | Ct value reported is for weakest sample pool
Centre National de Référence (CNR) Virus des Infections Respiratoires | Public Health | France | Paris | Confirmed COVID-19 patients | Not available | Nasopharyngeal swabs | Retrospective set | RT-PCR | High concentration pools: tested once; Low/neg concentration pools: tested in triplicate | Not available | Molecular | Manual PCR | Not available | iLAMP Novel-CoV19 Test Kit | ionebio | Not available | Not available | Nasopharyngeal swabs | Frozen RNA | Not available | Extraction NucleoSpin Dx Virus (Réf. Macherey Nagel 740895.50) | Not available | Not available | Not available | Not available | 29.2 | Overall | 9 | 8 | 1 | 3 | 1 | 37.5 | 100 | https://www.sfm-microbiologie.org/2020/04/03/covid-19/ | Not available | Ct value reported for lowest concentration pooled sample detected
Centre National de Référence (CNR) Virus des Infections Respiratoires | Public Health | France | Paris | Confirmed COVID-19 patients | Not available | Nasopharyngeal swabs | Retrospective set | RT-PCR | High concentration pools: tested once; Low/neg concentration pools: tested in triplicate | Not available | Molecular | Manual PCR | Not available | Bio-T kit SARS-CoV-2 | biosellal | Not available | Not available | Nasopharyngeal swabs | Frozen RNA | Not available | Extraction NucleoSpin Dx Virus (Réf. Macherey Nagel 740895.50) | Not available | Not available | Not available | Not available | 38 | Overall | 9 | 8 | 1 | 8 | 1 | 100 | 100 | https://www.sfm-microbiologie.org/2020/04/03/covid-19/ | Not available | Ct value reported is for weakest sample pool Amplification on ABI 7500 Fast Target: Gene E; ORF1ab
Centre National de Référence (CNR) Virus des Infections Respiratoires | Public Health | France | Paris | Confirmed COVID-19 patients | Not available | Nasopharyngeal swabs | Retrospective set | RT-PCR | High concentration pools: tested once; Low/neg concentration pools: tested in triplicate | Not available | Molecular | Manual PCR | Not available | Bio-T kit Covid-19 | biosellal | Not available | Not available | Nasopharyngeal swabs | Frozen RNA | Not available | Extraction NucleoSpin Dx Virus (Réf. Macherey Nagel 740895.50) | Not available | Not available | Not available | Not available | 37.4 | Overall | Not available | 8 | Not available | 8 | Not available | 100 |  | https://www.sfm-microbiologie.org/2020/04/03/covid-19/ | Not available | Ct value reported for most dilute pool Amplification on ABI 7500 Fast One target only
Centre National de Référence (CNR) Virus des Infections Respiratoires | Public Health | France | Paris | Confirmed COVID-19 patients | Not available | Nasopharyngeal swabs | Retrospective set | RT-PCR | High concentration pools: tested once; Low/neg concentration pools: tested in triplicate | Not available | Molecular | Manual PCR | Not available | Lyra SARS-CoV-2 Assay | Quidel | Not available | Not available | Nasopharyngeal swabs | Frozen RNA | Not available | Extraction NucleoSpin Dx Virus (Réf. Macherey Nagel 740895.50) | Not available | Not available | Not available | Not available | 27.7 | Overall | 9 | 8 | 1 | 6 | 0 | 75 | 0 | https://www.sfm-microbiologie.org/wp-content/uploads/2020/04/Rapport_Lyra-Quidel_VE200416.pdf | Not available | Ct value reported for lowest detectable concentration of pooled sample Amplification: ABI 7500 (also with LightCycler 480, but lower sensitivity)
Centre National de Référence (CNR) Virus des Infections Respiratoires | Public Health | France | Paris | Confirmed COVID-19 patients | Not available | Nasopharyngeal swabs | Retrospective set | RT-PCR | High concentration pools: tested once; Low/neg concentration pools: tested in triplicate | Not available | Molecular | Manual PCR | Not available | IDNCOV-2 | idSOLUTIONS | Not available | Not available | Nasopharyngeal swabs | Frozen RNA | Not available | Extraction NucleoSpin Dx Virus (Réf. Macherey Nagel 740895.50) | Not available | Not available | Not available | Not available | 35.5 | Overall | 9 | 8 | 1 | 8 | 1 | 100 | 100 | https://www.sfm-microbiologie.org/2020/04/03/covid-19/ | Not available | Ct value reported for lowest detectable concentration of pooled sample
Centre National de Référence (CNR) Virus des Infections Respiratoires | Public Health | France | Paris | Confirmed COVID-19 patients | Not available | Nasopharyngeal swabs | Retrospective set | RT-PCR | High concentration pools: tested once; Low/neg concentration pools: tested in triplicate | Not available | Molecular | Manual PCR | Not available | OPTI SARS-CoV-2 RT-PCR Test | OPTIMedical | Not available | Not available | Nasopharyngeal swabs | Frozen RNA | Not available | Extraction NucleoSpin Dx Virus (Réf. Macherey Nagel 740895.50) | Not available | Not available | Not available | Not available | 34.7 | Overall | 9 | 8 | 1 | 8 | 1 | 100 | 100 | https://www.sfm-microbiologie.org/2020/04/03/covid-19/ | Not available | Ct value reported for lowest detectable concentration of pooled sample Amplification on LightCycler480
Centre National de Référence (CNR) Virus des Infections Respiratoires | Public Health | France | Paris | Confirmed COVID-19 patients | Not available | Nasopharyngeal swabs | Retrospective set | RT-PCR | High concentration pools: tested once; Low/neg concentration pools: tested in triplicate | Not available | Molecular | Manual PCR | Not available | Corona Virus Disease 2019 (COVID-19) Nucleic Acid Detection Kit | EXCILONE SARL | Not available | Not available | Nasopharyngeal swabs | Frozen RNA | Not available | Extraction NucleoSpin Dx Virus (Réf. Macherey Nagel 740895.50) | Not available | Not available | Not available | Not available | 31.7 | Overall | 9 | 8 | 1 | 7 | 1 | 87.5 | 100 | https://www.sfm-microbiologie.org/2020/04/03/covid-19/ | Not available | Ct value reported for lowest detectable concentration of pooled sample Amplification by LightCycler 480 II outperformed ABI 7500; results here given for LightCycler 480 II
Centre National de Référence (CNR) Virus des Infections Respiratoires | Public Health | France | Paris | Confirmed COVID-19 patients | Not available | Nasopharyngeal swabs | Retrospective set | RT-PCR | High concentration pools: tested once; Low/neg concentration pools: tested in triplicate | Not available | Molecular | Manual PCR | Not available | EurobioPlex SARS-CoV-2 Multiplex | Eurobio Scientific | Not available | Not available | Nasopharyngeal swabs | Frozen RNA | Not available | Extraction NucleoSpin Dx Virus (Réf. Macherey Nagel 740895.50) | Not available | Not available | Not available | Not available | 35.5 | Overall | 8 | 7 | 1 | 7 | 1 | 100 | 100 | https://www.sfm-microbiologie.org/2020/04/03/covid-19/ | Not available | Not available
Centre National de Référence (CNR) Virus des Infections Respiratoires | Public Health | France | Paris | Confirmed COVID-19 patients | Not available | Nasopharyngeal swabs | Retrospective set | RT-PCR | High concentration pools: tested once; Low/neg concentration pools: tested in triplicate | Not available | Molecular | Manual PCR | Not available | SARS-CoV-2 Real Time PCR LAB-KIT™ | BIOMAXIMA S.A. | Not available | Not available | Nasopharyngeal swabs | Frozen RNA | Not available | Extraction NucleoSpin Dx Virus (Réf. Macherey Nagel 740895.50) | Not available | Not available | Not available | Not available | 38.1 | Overall | 8 | 7 | 1 | 7 | 1 | 100 | 100 | https://www.sfm-microbiologie.org/2020/04/03/covid-19/ | Not available | Ct value reported for lowest detectable concentration of pooled sample
Centre National de Référence (CNR) Virus des Infections Respiratoires | Public Health | France | Paris | Confirmed COVID-19 patients | Not available | Nasopharyngeal swabs | Retrospective set | RT-PCR | High concentration pools: tested once; Low/neg concentration pools: tested in triplicate | Not available | Molecular | Manual PCR | Not available | Q-Sens® 2019-nCoV Detection Kit | Cancer Rop Co., Ltd. | Not available | Not available | Nasopharyngeal swabs | Frozen RNA | Not available | Extraction NucleoSpin Dx Virus (Réf. Macherey Nagel 740895.50) | Not available | Not available | Not available | Not available | 36.7 | Overall | 9 | 8 | 1 | 8 | 1 | 100 | 100 | https://www.sfm-microbiologie.org/2020/04/03/covid-19/ | Not available | Test provider: Medlane France Amplification on ABI 7500
Centre National de Référence (CNR) Virus des Infections Respiratoires | Public Health | France | Paris | Confirmed COVID-19 patients | Not available | Nasopharyngeal swabs | Retrospective set | RT-PCR | High concentration pools: tested once; Low/neg concentration pools: tested in triplicate | Not available | Molecular | Manual PCR | Not available | 2019-Novel Coronavirus (2019-nCoV) RT-PCR Detection Kit | Shanghai Fosun Long March Medical Science Co.,Ltd. | Not available | Not available | Nasopharyngeal swabs | Frozen RNA | Not available | Extraction NucleoSpin Dx Virus (Réf. Macherey Nagel 740895.50) | Not available | Not available | Not available | Not available | 33 | Overall | 9 | 8 | 1 | 8 | 1 | 100 | 100 | https://www.sfm-microbiologie.org/2020/04/03/covid-19/ | Not available | Not available
Centre National de Référence (CNR) Virus des Infections Respiratoires | Public Health | France | Paris | Confirmed COVID-19 patients | Not available | Nasopharyngeal swabs | Retrospective set | RT-PCR | High concentration pools: tested once; Low/neg concentration pools: tested in triplicate | Not available | Molecular | Manual PCR | Not available | ID HRM SARS-COV-2 Triplex | ID.VET | Not available | Not available | Nasopharyngeal swabs | Frozen RNA | Not available | Extraction NucleoSpin Dx Virus (Réf. Macherey Nagel 740895.50) | Not available | Not available | Not available | Not available | Not available | Overall | 9 | 8 | 1 | 8 | 1 | 100 | 100 | https://www.sfm-microbiologie.org/2020/04/03/covid-19/ | Not available | Analysis of results through online software; no Ct values given, only binary outcome
Centre National de Référence (CNR) Virus des Infections Respiratoires | Public Health | France | Paris | Confirmed COVID-19 patients | Not available | Nasopharyngeal swabs | Retrospective set | RT-PCR | High concentration pools tested once Intermediate concentration/negative pools tested in triplicate Low concentration pools tested x5 | Not available | Molecular | Integrated system | Not available | Novadiag® COVID-19 + InfA/B (lab-based & near-POC) | Mobidiag | Not available | Not available | Nasopharyngeal swabs | Specimen without RNA extraction | Not available | Not available | Not available | Not available | Not available | Not available | Not available | Overall | 9 | 8 | 1 | 7 | 1 | 87.5 | 100 | https://www.sfm-microbiologie.org/2020/04/03/covid-19/ | Not available | Binary result, no Ct
Centre National de Référence (CNR) Virus des Infections Respiratoires | Public Health | France | Paris | Confirmed COVID-19 patients | Not available | Nasopharyngeal swabs | Retrospective set | RT-PCR | High concentration pools tested: x1 Intermediate concentration pools tested: x4 or x5 Negative pools tested: x2 | Not available | Molecular | Manual PCR | Not available | COVID-19 Multiplex Crystal Digital PCR detection kit | STILLA | Not available | Not available | Nasopharyngeal swabs | Frozen RNA | Not available | Extraction NucleoSpin Dx Virus (Réf. Macherey Nagel 740895.50) | Not available | Not available | Not available | Not available | Not available | Overall | 9 | 8 | 1 | 8 | 1 | 100 | 100 | https://www.sfm-microbiologie.org/2020/04/03/covid-19/ | Not available | Multiplex closed system amplification. Binary result only, no Ct provided. Method is quantitative
Centre National de Référence (CNR) Virus des Infections Respiratoires | Public Health | France | Paris | Confirmed COVID-19 patients | Not available | Nasopharyngeal swabs | Retrospective set | RT-PCR | High concentration pools: tested once; Low/neg concentration pools: tested in triplicate | Not available | Molecular | Manual isothermal amplification | Not available | SmartAmp Covid-19 (Lamp system) | EMG / MSE | Not available | Not available | Nasopharyngeal swabs | Fresh RNA extracted | Not available | LifeCase RNA extraction | Not available | Not available | Not available | Not available | Not available | Overall | 9 | 8 | 1 | 7 | 1 | 87.5 | 100 | https://www.sfm-microbiologie.org/2020/04/03/covid-19/ | Not available | Amplification compared between: LifeCase and LightCycler 480: same performance
Centre National de Référence (CNR) des Virus des Infections Respiratoires Institut des Agents infectieux, Centre de Biologie et Pathologie Nord | Public Health | France | Lyon | Confirmed COVID-19 patients | Not available | Nasopharyngeal swabs | Retrospective set | RT-PCR | Not available | Not available | Molecular | Manual PCR | Not available | VIASURE SARS-CoV-2 Real Time PCR Detection Kit | CerTest Biotec, S.L. | Not available | Not available | Nasopharyngeal swabs | Frozen RNA | Not available | EMAG® bioMérieux | Not available | Not available | Not available | Not available | 44.8 | Overall | 20 | 18 | 2 | 17 | 2 | 94.4 | 100 | https://www.sfm-microbiologie.org/2020/04/03/covid-19/ | Not available | Ct value is highest value across either loci detected positive.
Centre National de Référence (CNR) des Virus des Infections Respiratoires Centre de Biologie et Pathologie Nord | Public Health | France | Lyon | Confirmed COVID-19 patients | Not available | Nasopharyngeal swabs | Retrospective set | RT-PCR | Triplicates | Not available | Molecular | Integrated system | Not available | BioFire COVID-19 Test | bioMérieux SA | Not available | Not available | Nasopharyngeal swabs | Specimen without RNA extraction | Not available | Not available | Not available | Not available | Not available | Not available | Not available | Overall | 15 | 11 | 4 | 11 | 4 | 100 | 100 | https://www.sfm-microbiologie.org/2020/04/03/covid-19/ | Not available | Positive results corresponded to Ct range 19-35.9 on reference method
Centre National de Référence des virus des infections respiratoires Centre de Biologie et Pathologie Nord | Public Health | France | Lyon | Confirmed COVID-19 patients | Not available | Nasopharyngeal swabs; Tracheal aspirate | Retrospective set | RT-PCR | Not available | Not available | Molecular | Manual PCR | Not available | Bosphore Novel Coronavirus (2019-nCoV) Detection Kit | Anatolia Geneworks | ABCOW5 | Not available | Nasopharyngeal swabs; Tracheal aspirate | Frozen RNA | Not available | EMAG® bioMérieux | Not available | Not available | Not available | Not available | 33 | Overall | 13 | 10 | 3 | 8 | 3 | 80 | 100 | https://www.sfm-microbiologie.org/2020/04/03/covid-19/ | Not available | Some samples needed repeating; results given are final outcomes from Table 1 & 2 Ct given is highest value for a positively detected sample
Laboratoire associé au Centre National de Référence des virus des infections respiratoires Centre de Biologie et Pathologie Nord | Public Health | France | Lyon | Confirmed COVID-19 patients | Not available | Nasopharyngeal swabs | Retrospective set | RT-PCR | Not available | Not available | Molecular | Manual PCR | Not available | UltraGene Combo2Screen SARS-CoV-2 Assay | Advanced Biological Laboratories SA | 139A100 | Not available | Nasopharyngeal swabs | Frozen RNA | Not available | EMAG® bioMérieux | Not available | Not available | Not available | Not available | 35.7 | Overall | 12 | 10 | 2 | 8 | 2 | 80 | 100 | https://www.sfm-microbiologie.org/2020/04/03/covid-19/ | Not available | Ct value reported was highest for a positively detected sample
Johns Hopkins University; LAC+USC Medical Center, University of Southern California; Manchester University NHS Foundation Trust; Mondor Hospital; New York City Dept. Health and Mental Hygiene; Niguarda Hospital; Newark University Hospital | Not available | United States of America (USA) | Not available | Confirmed COVID-19 patients | Not available | Nasopharyngeal swabs; Oropharyngeal swabs; Tracheal aspirate | Prospectively collected set | RT-PCR | Not available | Not available | Molecular | Integrated system | Not available | Xpert Xpress SARS-CoV-2 | Cepheid | Not available | Not available | Nasopharyngeal swabs; Oropharyngeal swabs; Tracheal aspirate | Specimen without RNA extraction | Not available | Not available | 0 | PFU/mL | Not available | Not available | 37.1 | Overall | 481 | 220 | 261 | 219 | 250 | 99.5 | 95.8 | https://jcm.asm.org/content/jcm/early/2020/04/30/JCM.00926-20.full.pdf | 10.1128/JCM.00926-20 | Note, manufacturer is among authors. Target: Envelope (E), nucleocapsid (N2) genes Specimen types differed among some sites. Specimen sources included NPS (339), combined NPS/OPS (97), TA (30), and OPS (15).
Northwell Health Laboratories | Not available | United States of America (USA) | New York | Symptomatic patients | Not available | Nasopharyngeal swabs | Retrospective set | RT-PCR | Not available | Not available | Molecular | Integrated system | Not available | Xpert Xpress SARS-CoV-2 | Cepheid | Not available | Not available | Nasopharyngeal swabs | Specimen without RNA extraction | Not available | Not available | 100 | copy / mL | Not available | Not available | Not available | Overall | 108 | 58 | 50 | 57 | 50 | 98.3 | 100 | https://jcm.asm.org/content/jcm/early/2020/04/23/JCM.00783-20.full.pdf | 10.1128/JCM.00783-20 | Not available
University of Washington Medical Center laboratories | Not available | United States of America (USA) | Washington | Confirmed COVID-19 patients | Not available | Nasopharyngeal swabs | Retrospective set | RT-PCR | No replicates | Not available | Molecular | Integrated system | Not available | Panther Fusion SARS-CoV-2 assay | Hologic | Not available | Not available | Nasopharyngeal swabs | Specimen without RNA extraction | Not available | Not available | Not available | Not available | Not available | Not available | 37 | Overall | 46 | 17 | 29 | 16 | 29 | 94.1 | 100 | https://jcm.asm.org/content/jcm/early/2020/04/27/JCM.00821-20.full.pdf | 10.1128/JCM.00821-20 | Research use only (RUO) assay
Bichat Claude Bernard Teaching Hospital | Hospital | France | Paris | Symptomatic patients | Not available | Nasopharyngeal swabs | Prospectively collected set | RT-PCR | Not available | Not available | Molecular | Integrated system | Not available | QIAstat-Dx Respiratory Panel 2019-nCoV | QIAGEN GmbH | Not available | Not available | Nasopharyngeal swabs; Lower respiratory specimen | Specimen without RNA extraction | Not available | Not available | 1000 | copy / mL | Not available | Not available | Not available | Overall | 69 | 40 | 29 | 40 | 27 |  |  | https://jcm.asm.org/content/jcm/early/2020/04/24/JCM.00630-20.full.pdf | 10.1128/JCM.00630-20 | Not available
University of Washington Medical Center laboratories | Not available | United States of America (USA) | Washington | Confirmed COVID-19 patients | Not available | Nasopharyngeal swabs | Retrospective set | RT-PCR | No replicates | Not available | Molecular | Integrated system | Not available | Panther Fusion SARS-CoV-2 assay | Hologic | Not available | Not available | Nasopharyngeal swabs | Specimen without RNA extraction | Not available | Not available | Not available | Not available | Not available | Not available | Not available | Overall | Not available | 12 | Not available | 11 | Not available | 91.7 |  | https://jcm.asm.org/content/jcm/early/2020/04/27/JCM.00821-20.full.pdf | 10.1128/JCM.00821-20 | Not available
University of Washington Medical Center laboratories | Not available | United States of America (USA) | Washington | Confirmed COVID-19 patients | Not available | Nasopharyngeal swabs | Retrospective set | RT-PCR | Not available | Not available | Molecular | Integrated system | Not available | Simplexa™ COVID-19 Direct RT-PCR Kit | DiaSorin Molecular, LLC | Not available | Not available | Nasopharyngeal swabs | Specimen without RNA extraction | Not available | Not available | Not available | Not available | Not available | Not available | Not available | Overall | Not available | 13 | Not available | 12 | Not available | 92.3 |  | https://jcm.asm.org/content/jcm/early/2020/04/27/JCM.00821-20.full.pdf | 10.1128/JCM.00821-20 | Not available
University of Washington Medical Center laboratories | Not available | United States of America (USA) | Washington | Confirmed COVID-19 patients | Not available | Nasopharyngeal swabs | Retrospective set | RT-PCR | No replicates | Not available | Molecular | Integrated system | Not available | Panther Fusion SARS-CoV-2 assay | Hologic | Not available | Not available | Nasopharyngeal swabs | Specimen without RNA extraction | Not available | Not available | Not available | Not available | Not available | Not available | 37 | Overall | 20 | 10 | 10 | 10 | 9 | 100 | 90 | https://jcm.asm.org/content/jcm/early/2020/04/27/JCM.00821-20.full.pdf | 10.1128/JCM.00821-20 | Emergency Use Authorization (EUA) assay
University of Washington Medical Center laboratories | Not available | United States of America (USA) | Washington | Confirmed COVID-19 patients | Not available | Nasopharyngeal swabs | Retrospective set | RT-PCR | No replicates | Not available | Molecular | Integrated system | Not available | cobas® SARS-CoV-2 (for use on the cobas® 6800/8800 Systems) | Roche Molecular Diagnostics | Not available | Not available | Nasopharyngeal swabs | Specimen without RNA extraction | Not available | Not available | Not available | Not available | Not available | Not available | Not available | Overall | 26 | 13 | 13 | 12 | 13 | 92.3 | 100 | https://jcm.asm.org/content/jcm/early/2020/04/27/JCM.00821-20.full.pdf | 10.1128/JCM.00821-20 | Not available
University of Washington Medical Center laboratories | Not available | United States of America (USA) | Washington | Confirmed COVID-19 patients | Not available | Nasopharyngeal swabs | Retrospective set | RT-PCR | No replicates | Not available | Molecular | Integrated system | Not available | Simplexa™ COVID-19 Direct RT-PCR Kit | DiaSorin Molecular, LLC | Not available | Not available | Nasopharyngeal swabs | Specimen without RNA extraction | Not available | Not available | Not available | Not available | Not available | Not available | Not available | Overall | 19 | 11 | 8 | 11 | 8 | 100 | 100 | https://jcm.asm.org/content/jcm/early/2020/04/27/JCM.00821-20.full.pdf | 10.1128/JCM.00821-20 | Not available
Northwell Health Laboratories | Not available | United States of America (USA) | New York | Symptomatic patients | Not available | Nasopharyngeal swabs | Retrospective set | RT-PCR | Not available | Not available | Molecular | Integrated system | Not available | ID NOW COVID-19 | Abbott Diagnostics Inc. | Not available | Not available | Nasopharyngeal swabs | Specimen without RNA extraction | Not available | Not available | 20000 | copy / mL | Not available | Not available | Not available | Overall | 107 | 57 | 50 | 50 | 50 | 87.7 | 100 | https://jcm.asm.org/content/jcm/early/2020/04/23/JCM.00783-20.full.pdf | 10.1128/JCM.00783-20 | Not available
University of Washington Medical Center laboratories | Not available | United States of America (USA) | Washington | Confirmed COVID-19 patients | Not available | Nasopharyngeal swabs | Retrospective set | RT-PCR | No replicates | Not available | Molecular | Integrated system | Not available | cobas® SARS-CoV-2 (for use on the cobas® 6800/8800 Systems) | Roche Molecular Diagnostics | Not available | Not available | Nasopharyngeal swabs | Specimen without RNA extraction | Not available | Not available | Not available | Not available | Not available | Not available | Not available | Overall | 40 | 20 | 20 | 20 | 19 | 100 | 95 | https://jcm.asm.org/content/jcm/early/2020/04/27/JCM.00821-20.full.pdf | 10.1128/JCM.00821-20 | Not available
Northwell Health Laboratories | Not available | United States of America (USA) | New York | Symptomatic patients | Not available | Nasopharyngeal swabs | Retrospective set | RT-PCR | Not available | Not available | Molecular | Integrated system | Not available | ePlex® SARS-CoV-2 Test | GenMark Diagnostics | Not available | Not available | Nasopharyngeal swabs | Specimen without RNA extraction | Not available | Not available | 1000 | copy / mL | Not available | Not available | Not available | Overall | 108 | 58 | 50 | 53 | 50 | 91.4 | 100 | https://jcm.asm.org/content/jcm/early/2020/04/23/JCM.00783-20.full.pdf | 10.1128/JCM.00783-20 | Not available
University of Washington Medical Center laboratories | Not available | United States of America (USA) | Washington | Confirmed COVID-19 patients | Not available | Nasopharyngeal swabs | Retrospective set | RT-PCR | Not available | Not available | Molecular | Integrated system | Not available | Xpert Xpress SARS-CoV-2 | Cepheid | Not available | Not available | Nasopharyngeal swabs | Specimen without RNA extraction | Not available | Not available | Not available | Not available | Not available | Not available | Not available | Overall | 26 | 13 | 13 | 13 | 13 | 100 | 100 | https://jcm.asm.org/content/jcm/early/2020/04/27/JCM.00821-20.full.pdf | 10.1128/JCM.00821-20 | Not available
General Hospital of Central Theater Command of PLA | Hospital | China | Wuhan | Symptomatic patients | Not available | Pharyngeal swabs | Prospectively collected set | Clinical examination; RT-PCR | Not available | Not available | Molecular | Manual PCR | Not available | Novel Coronavirus (2019-nCoV) Real Time Multiplex RT-PCR kit | DAAN Gene Co., Ltd of Sun Yat-sen University | Not available | Not available | Pharyngeal swabs | Fresh RNA extracted | Not available | Shuoshi, Shanghai, China | Not available | Not available | Not available | Not available | Not available | Overall | Not available | 238 | Not available | 153 | Not available | 64.3 |  | https://www.medrxiv.org/content/10.1101/2020.03.06.20031856v1.full.pdf | 10.1101/2020.03.06.20031856 | Not available
Centre Intégré Universitaire de Santé et Services Sociaux de la Mauricie et Centre du Québec | Public Health | Canada | Trois-Rivieres | Symptomatic patients | Not available | Nasopharyngeal swabs; Oropharyngeal swabs | Prospectively collected set | Not available | No replicates | Not available | Molecular | Manual PCR | Not available | RealStar® SARS-CoV-2 RT-PCR Kit | altona Diagnostics | Not available | Not available | Nasopharyngeal swabs; Oropharyngeal swabs | Frozen RNA | Not available | Abbott mSample Preparation Systems DNA kit on m2000sp instrument (Abbott) | Not available | Not available | Not available | Not available | Not available | Overall | 88 | 65 | 23 | 65 | 23 | 100 | 100 | https://www.ncbi.nlm.nih.gov/pmc/articles/PMC7204723/ | 10.1016/j.jcv.2020.104423 | Not available
Centre Intégré Universitaire de Santé et Services Sociaux de la Mauricie et Centre du Québec | Public Health | Canada | Trois-Rivieres | Symptomatic patients | Not available | Nasopharyngeal swabs; Oropharyngeal swabs | Prospectively collected set | Not available | Not available | Not available | Molecular | Manual PCR | Not available | Allplex 2019-nCoV assay | Seegene, Inc. | Not available | Not available | Nasopharyngeal swabs; Oropharyngeal swabs | Frozen RNA | Not available | Not available | Not available | Not available | Not available | Not available | 40 | Overall | 88 | 65 | 23 | 64 | 23 | 98.5 | 100 | https://www.ncbi.nlm.nih.gov/pmc/articles/PMC7204723/ | 10.1016/j.jcv.2020.104423 | Not available
National Institute for Infectious Diseases ‘Lazzaro Spallanzani’ IRCC; Not available | Public Health | Italy | Rome | Symptomatic patients | Not available | Nasopharyngeal swabs; Nasal swabs | Prospectively collected set | RT-PCR | Not available | Not available | Molecular | Integrated system | Not available | Simplexa™ COVID-19 Direct RT-PCR Kit | DiaSorin Molecular, LLC | Not available | Not available | Nasopharyngeal swabs | Specimen without RNA extraction | Not available | Not available | Not available | Not available | Not available | Not available | Not available | Overall | 278 | 99 | 179 | 99 | 171 |  |  | https://reader.elsevier.com/reader/sd/pii/S138665322030158X?token=E8321EF3F282976013F76FA5A1542161F66A72DA794A2EABD1A2107016BE331A2D5214F28B1C6EED052DA93074681304 | 10.1016/j.jcv.2020.104416 | LOD for both target genes: 3.2 (CI: 2.9 to 3.8) log10 copy/mL
Department of Microbiology, Monash Health, Clayton, VIC | Hospital | Australia | Clayton | Symptomatic patients | Not available | Nasopharyngeal swabs; Oropharyngeal swabs | Prospectively collected set | RT-PCR | Not available | Not available | Molecular | Integrated system | Not available | SARS-CoV-2, Influenza and RSV 8-well | AusDiagnostics | 20081 | Not available | Nasopharyngeal swabs; Oropharyngeal swabs; Sputum; Tracheal aspirate | Fresh RNA extracted | Not available | AusDiagnostics MT-Prep extraction system | Not available | Not available | Not available | Not available | Not available | Overall | 187 | 127 | 60 | 118 | 59 | 92.9 | 98.3 | https://www.ncbi.nlm.nih.gov/pmc/articles/PMC7236671/ | 10.1016/j.jcv.2020.104448 | Sample transportation in UTM medium
Public Health England (PHE) Cambridge laboratory | Public Health | Not available | Cambridge | Symptomatic patients | Not available | Nasopharyngeal swabs; Oropharyngeal swabs | Retrospective set | RT-PCR | Not available | Blinded to reference results | Molecular | Integrated system | Not available | SAMBA II COVID-19 Test (near-POC NAT) | Diagnostics for the Real World Ltd | Not available | Not available | Nasopharyngeal swabs; Oropharyngeal swabs | Specimen without RNA extraction | Not available | Not available | 250 | copy / mL | Not available | Not available | 36 | Overall | 172 | 91 | 81 | 90 | 81 | 98.9 | 100 | https://www.medrxiv.org/content/10.1101/2020.05.24.20100990v2.full.pdf | 10.1101/2020.05.24.20100990 | Not available
Department of Microbiology, Queen Mary Hospital | Hospital | China - Hong Kong / Macau | Hong Kong | Symptomatic patients | Not available | Nasopharyngeal swabs | Retrospective set | RT-PCR | Not available | Not available | Molecular | Not available | Not available | NxTAG CoV Extended Panel | Luminex Corporation | Not available | Not available | Nasopharyngeal swabs | Frozen RNA | Not available | Not available | 200 | copy / mL | Not available | Not available | Not available | Overall | 214 | 91 | 123 | 89 | 123 | 97.8 | 100 | https://jcm.asm.org/content/jcm/early/2020/05/29/JCM.00936-20.full.pdf | 10.1128/JCM.00936-20 | Not available
Department of Pathology and Clinical Laboratory Services, University of Texas Southwestern Medical Center | Hospital | United States of America (USA) | Dallas | Symptomatic patients | Not available | Nasopharyngeal swabs | Retrospective set | RT-PCR | No replicates | Not available | Molecular | Integrated system | Not available | ID NOW COVID-19 | Abbott Diagnostics Inc. | Not available | Not available | Dry nasal swabs | Specimen without RNA extraction | Not available | Not available | Not available | Not available | Not available | Not available | Not available | Overall | 579 | 33 | 546 | 26 | 546 | 78.8 | 100 | https://onlinelibrary.wiley.com/doi/epdf/10.1111/acem.14039 | 10.1111/acem.14039 | Not available
University Medical Center Hamburg-Eppendorf (UKE), Institute of Medical Microbiology, Virology and Hygiene | Hospital | Germany | Hamburg | Confirmed COVID-19 patients | Not available | Nasopharyngeal swabs; Oropharyngeal swabs | Retrospective set | RT-PCR | Not available | Not available | Molecular | Integrated system | Not available | cobas® SARS-CoV-2 (for use on the cobas® 6800/8800 Systems) | Roche Molecular Diagnostics | Not available | Not available | Nasopharyngeal swabs; Oropharyngeal swabs | Specimen without RNA extraction | Not available | Not available | Not available | Not available | Not available | Not available | Not available | Overall | 180 | 93 | 87 | 89 | 86 | 95.7 | 98.9 | https://www.medrxiv.org/content/10.1101/2020.05.28.20115469v1.full.pdf | 10.1101/2020.05.28.20115469 | Not available
Department of Microbiology, Royal Melbourne Hospital | Hospital | Australia | Melbourne | Symptomatic patients | Not available | Bronchoalveolar lavage; Nasopharyngeal swabs; Oropharyngeal swabs; Sputum; Tracheal aspirate | Prospectively collected set | RT-PCR | Not available | Not available | Molecular | Integrated system | Not available | SARS-CoV-2, Influenza and RSV 8-well | AusDiagnostics | Not available | Not available | Bronchoalveolar lavage; Nasopharyngeal swabs; Oropharyngeal swabs; Sputum; Tracheal aspirate | Fresh RNA extracted | Not available | Viral RNA mini kit (QIAGEN, Hilden, Germany) | Not available | Not available | Not available | Not available | Not available | Overall | 675 | 53 | 622 | 53 | 621 | 100 | 99.8 | https://www.medrxiv.org/content/10.1101/2020.06.03.20117267v1.full.pdf | 10.1101/2020.06.03.20117267 | Not available
Wuhan Tongji Hospital; Wuhan Pulmonary 47 Hospital; Wuhan No. 1 Hospital | Hospital | China | Wuhan | Not available | Mild; Moderate; Severe; Critical | Oropharyngeal swabs | Retrospective set | RT-PCR | Triplicates | Not available | Molecular | Integrated system | Not available | Xpert Xpress SARS-CoV-2 | Cepheid | Not available | Not available | Oropharyngeal swabs | Specimen without RNA extraction | Not available | Not available | Not available | Not available | Not available | Not available | Not available | Overall | 285 | 153 | 132 | 147 | 127 | 96.1 | 96.2 | https://jcm.asm.org/content/jcm/early/2020/06/05/JCM.01288-20.full.pdf | 10.1128/JCM.01288-20 | Not available
University of Washington Retrovirology Laboratory at Harborview Medical Center | Hospital | United States of America (USA) | Seattle | Confirmed COVID-19 patients | Not available | Nasopharyngeal swabs; Nasal swabs | Retrospective set | RT-PCR | Not available | Blinded to reference results | Molecular | Integrated system | Not available | Abbott RealTime SARS-CoV-2 EUA test | Abbott Molecular Inc. | Not available | Not available | Nasopharyngeal swabs; Nasal swabs | Specimen without RNA extraction | Not available | Not available | 100 | copy / mL | Not available | Not available | Not available | Overall | 60 | 30 | 30 | 28 | 30 | 93.3 | 100 | https://www.sciencedirect.com/science/article/pii/S138665322030216X?via%3Dihub | 10.1016/j.jcv.2020.104474 | LoD was <100 copies/mL in 19/20 samples
Clinical Virology Laboratory, Stanford Health Care | Not available | United States of America (USA) | Stanford | Not available | Asymptomatic; Mild; Moderate; Severe | Nasopharyngeal swabs | Prospectively collected set | RT-PCR | Not available | Not available | Molecular | Integrated system | Not available | Accula SARS-CoV-2 POCT | Mesa Biotech Inc. | Not available | Not available | Nasopharyngeal swabs | Specimen without RNA extraction | Not available | Not available | Not available | Not available | Not available | Not available | Not available | Overall | 100 | 50 | 50 | 34 | 50 | 68 | 100 | https://jcm.asm.org/content/early/2020/05/25/JCM.01072-20 | 10.1128/JCM.01072-20 | Not available
Rush University Medical Center or Rush Oak Park Hospital | Hospital | United States of America (USA) | Chicago | Symptomatic patients | Mild; Moderate; Severe; Critical | Nasopharyngeal swabs | Prospectively collected set | RT-PCR | Not available | Blinded to reference results | Molecular | Integrated system | Not available | Abbott RealTime SARS-CoV-2 EUA test | Abbott Molecular Inc. | Not available | Not available | Nasopharyngeal swabs | Specimen without RNA extraction | Not available | Not available | Not available | Not available | Not available | Not available | Not available | Overall | 198 | 119 | 79 | 119 | 73 | 100 | 92.4 | https://jcm.asm.org/content/jcm/early/2020/05/25/JCM.00938-20.full.pdf | 10.1128/JCM.00938-20 | Not available
Rush University Medical Center or Rush Oak Park Hospital | Hospital | United States of America (USA) | Chicago | Symptomatic patients | Mild; Moderate; Severe; Critical | Nasopharyngeal swabs | Prospectively collected set | RT-PCR | Not available | Blinded to reference results | Molecular | Integrated system | Not available | ID NOW COVID-19 | Abbott Diagnostics Inc. | Not available | Not available | Nasopharyngeal swabs | Specimen without RNA extraction | Not available | Not available | Not available | Not available | Not available | Not available | Not available | Overall | 198 | 119 | 79 | 94 | 79 | 79 | 100 | https://jcm.asm.org/content/jcm/early/2020/05/25/JCM.00938-20.full.pdf | 10.1128/JCM.00938-20 | Not available